## Sheet1
A) LA-MC-ICP-MS U-Th-Pb data report PEDOGENIC CARBONATES
Laboratory & Sample Preparation
Laboratory name | FIERCE, Frankfurt Isotope & Element Research Center Goethe Univesität, Frankfurt am Main
Sample type/mineral | Carbonate
Sample preparation | standard thin section
Imaging | Petrographic microscope & 2400 dpi digital scan
Laser ablation system
Make, Model & type | RESOlution ArF excimer laser (COMpex Pro 102)
Ablation cell | Two-volume ablation cell (Laurin Technic S155)
Laser wavelength | 193 nm
Pulse width | 20 ns
Fluence | 2 J/cm2
Repetition rate | 8 Hz
Pre-ablation | 20 pulses (same parameters as main ablation)
Ablation duration | 18 s
Ablation rate | ~ 0.6 μm/s
Spot shape & size | Round, 104 mm (diameter)
Sampling mode | Static spot ablation
Gasses | Sample cell: He. Funnel: He + Ar. Tubbing: He + Ar + N
Gas flows | He (300 ml/min), Ar (1000 ml/min), N (8 ml/min).
ICP-MS Instrument
Make, Model & type | ThermoScientific Neptune+ MC-ICP-MS
Sample introduction | Ablation aerosol
RF power | 1300 W
Detection system | Static detection with 207Pb and 206Pb on full-sized secondary electron multiplier and 232Th and 238U on Faraday collector, amplifier with 1013 O resitors. Ion counter yield and faraday/SEM gain calibrated with 235U solution (5 mV 235U on Faraday cups)
Masses measured | (202, 204), 206, 207, 232, 238
Dwell times | 232 ms
Total time per run | 232 ms
Number of runs/total time | 135 / 35.4 s
Acquisition mode | Trigger from laser (20 s after pre-ablation), background: 18 s, ablation: 18 s
Dead time | 29 ns
Data Processing
Gas blank | 18 s on-peak zero subtracted.
Calibration strategy | NIST SRM-614 used as primary standard, WC-1 for carbonate matrix offset determination, and ASH-15D, CalGr, and B6 as secondary reference materials (validation).
Reference Material (RM) information | Soda-lime glass NIST SRM-614, WC-1 (Roberts et al., 2017), ASH-15D (Vaks et al., 2003), B6 (Pagel et al. 2018), Calgrun (in-house calcite RM)
Data processing / LIEF correction | In-house VBA spreadsheet program modified after Gerdes and Zeh (2006, 2009). No LIEF correction for Pb/U.
Mass discrimination | 207Pb/206Pb (0.4%) and 206Pb/238U (5%) normalised to primary standard
Common-Pb correction | No common-Pb correction applied to the data.
Uncertainty level & propagation | Ages are quoted at 2s absolute, propagation by quadratic addition.
Quality control / Validation | Sequence 1: WC-1: 254.18 ±0.99 (2s, MSWD = 1.3, n = 8) ASH-15D: 3.01 ± 0.31 Ma (2s, MSWD = 0.41, n = 8) Sequence 2: WC-1: 254.4 ± 1.4 Ma (2s, MSWD = 1.86, n = 14) ASH-15D: 2.91 ± 0.14 Ma (2s, MSWD = 0.41, n = 19) B6: 42.12 ± 0.31 Ma (2s, MSWD = 1.01, n = 14) and CalGr: 20.13 ± 0.11 Ma (2s, MSWD = 1.55, n = 14)
B) LA-MC-ICP-MS U-Th-Pb data report BOVID TOOTH
Laboratory & Sample Preparation
Laboratory name | FIERCE, Frankfurt Isotope & Element Research Center Goethe Univesität, Frankfurt am Main
Sample type/mineral | Phosphate
Sample preparation | 25 mm polished resin mounts
Imaging | Petrographic microscope & 2400 dpi digital scan
Laser ablation system
Make, Model & type | RESOlution ArF excimer laser (COMpex Pro 102)
Ablation cell | Two-volume ablation cell (Laurin Technic S155)
Laser wavelength | 193 nm
Pulse width | 20 ns
Fluence | 2 J/cm2
Repetition rate | 6 Hz
Pre-ablation | 2 pulses (same parameters as main ablation)
Ablation duration | 18 s
Ablation rate | ~ 0.6 μm/s
Spot shape & size | round, 50 mm (diameter)
Sampling mode | Static spot ablation
Gasses | Sample cell: He. Funnel: He + Ar. Tubbing: He + Ar + N
Gas flows | He (300 ml/min), Ar (1100 ml/min), N (5 ml/min).
ICP-MS Instrument
Make, Model & type | ThermoScientific ElementXr sector field ICP-MS
Sample introduction | Ablation aerosol
RF power | 1300 W
Detection system | Secondary electron multiplier (with conversion dynode at -8kV). Simultaneous analogue and counting (pulse) modes of detection (conversion factors calculated per mass and applied offline). Magnetic field fixed. Detection by peak jumping with electrostatic analyzer.
Masses measured | 206, 207, 208, 232, 238
Dwell times | 206: 6.4 ms, 207: 7.5 ms, 208: 3.0 ms, 232: 2.0 ms, 238: 4.6 ms
Samples per peak/integration type | 4 for all masses/average
Total time per run | 99 ms
Number of runs/total time | 370 / 36.6 s
Acquisition mode | Trigger from laser (20 s after pre-ablation), background: 18 s, ablation: 18 s
Dead time | 29 ns
Gas blank | 18 s on-peak zero subtracted.
Calibration strategy | NIST SRM-612 as primary RM, Durango apatite for matrix offset on 206Pb/238U (2.4%), and Bra Ap as validation RM.
Reference Material (RM) information | Soda-lime glass NIST SRM-612, Durango apatite (McDowell et al., 2005), Bra Ap (in-house apatite RM)
Data processing / LIEF correction | In-house VBA spreadsheet program (Gerdes and Zeh, 2006, 2009).
Mass discrimination | 207Pb/206Pb (0.2%) and 206Pb/238U (10%) normalised to primary standard
Common-Pb correction | No common-Pb correction applied to the data.
Uncertainty level & propagation | Uncertainties are quoted at 2δ absolute and are propagated by quadratic addition of the within run precision (SD of the mean of ratios in log-ratio space), counting statistics, background, common Pb correction (if applicable) and the excess of scatter (calculated from the primary RM). In addition, an excess of variance of 1.4 % (1δ), calculated from the offset RM, was added quadratically to the 206Pb/238U ratios. Systematic uncertainties are reported as an expanded uncertainty, considering long term reproducibility (1.0%, 2δ) and decay constant uncertainties.